Thank you very much for you review and valuable suggestions. Below you can 
find addressed all points from the review: 

1. What is the optical bandwidth (wavelength range, central wavelength) 
used 
for the imaging? 

The system axial resolution of 8 mm in tissue is defined by 67 nm spectral 
bandwidth of the polygon based laser source centered at 1310 nm. 

How deeply can this light penetrate into the imaged tissue? 

Spectral properties of the imaging light beam allows deep tissue 
penetration up to the muscularis propria of the human esophagus (about 1.5 
mm). 
2. Information regarding the imaging core optics (page 2) is somewhat 
vague.What do authors mean by "a beam size of 30um FWHM"? Is this the spot 
diameter of the focused beam, i.e. the lateral imaging resolution of the 
system? Is this really the FWHM diameter or rather 1/e^2 diameter? Please 
clarify this. 

The light is focused by the GRIN lens to a spot diameter of 30 mm 
measured at FWHM. As shown at Figure 2, the light is directed to the side 
by a 
right angle prism assembled at the tip of the imaging core [11]. The beam 
is focused 
at 12.5 mm from the centre of the optical core and a its confocal 
parameter equals 
3.28 mm. 

3. Information about the size (diameter, length) of the distal 
opticalassembly is missing. 

.....which we term the imaging core, comprising an optical fiber threaded 
into 1 mm long glass ferrule, a 2.5 mm long spacer, a 2.01 mm long 
gradient index (GRIN) lens, and a prism with aluminium coating on 
hypotenuse. The outer diameter of the optical imaging core is 1 mm. 
4. Labeling different parts of the optical probe in the photograph in 
Fig.2would help the reader to associate them with the schematic drawing. 

Description added to the Fig.2 

5. Image size indication (e.g. scale bars) is missing in figures 4, 5, 6d, 
therefore it is impossible to judge what are the sizes of the tissue 
images. 

Scale bars added to Fig.4, 5, 6d 

6. Please indicate the wires of the "basket" in the image of the swine 
ileum 
(figure 6d). 

Centering wires are pointed out with arrows at Fig6d 

7. In the third sentence of the paper: "Light returned from the sample and 
reference are combined and detected". It seems that word "arm" or "channel" 
is missing 

Missing word "arm" was added: 
Light returned from the sample and reference arms are combined and 
detected. 

Thanks you very much for you time and please let me know if you have any 
further questions. 

Best regards 
Michalina Gora